# V.Epshteyn, P.Shatalov     tests_2003 Note2, 29-Aug-2003

Last update: 14-Sep-2003 (with P.Gorbunov)

## Some results of 2003 tests

1. Introduction

We present a summary of the analysis we did for selected TB’2003 runs taken in summer 2003 and imported to ITEP in July  (Table 1). Our aim was to review the main features of the TB data and evaluate the observed energy resolution for electrons and pions, in the 40-200 GeV/c beam momentum range.

| Table 1: Data sample used in the present analysis |
| --- |
| beam type / position | P, GeV/c | runs # | number of events |
| electrons / 4H | 40 | 2845, 2846, 2847 | 36000 |
|  | 60 | 3328, 3329 | 24000 |
|  | 100 | 2687, 2688 | 24000 |
|  | 200 | 2967, 2968 | 24000 |
| electrons / 4L | 60 | 3242, 3243 | 24000 |
|  | 100 | 2595, 2596 | 24000 |
|  | 200 | 3223, 3224 | 24000 |
| pions / 4H | 60 | 3294 | 12000 |
|  | 100 | 2649 | 12000 |
|  | 200 | 3043 | 12000 |
| pions / 4L | 60 | 3259 | 12000 |
|  | 100 | 2613 | 12000 |
|  | 200 | 3193 | 12000 |

Some aspects of the signal and energy reconstruction remained beyond the scope of  the present analysis: a detailed signal shape study, the channel equalization, cell clusters etc. On the other hand, we made an attempt to optimize the module intercalibration for the hadronic energy reconstruction and studied the effect of the signal filtering method, as well as the event selection cuts, on the calorimeter response.

2. Signal amplitude

Let Ai  be the signal value corresponding to the sample i=0,1…6.   Three methods to reconstruct the signal amplitude A (to filter the signal) are considered (Fig.1):

- raw : A=max(Ai).

- parabola:  A is the maximum of a parabola fit to A2,3,4

- spline – A is the maximum the cubic spline drawn through all the samples.

A certain contribution to the energy resolution comes from the signal quantization noise. Its dependence on the filtering method is illustrated by a reconstruction of the simulated shaper signal  of a random phase:  the relative rms spread of the reconstructed amplitude is 6.8%, 3.7% and 2.6% for the “raw”, “parabola” and the “spline” methods, respectively. The better we approximate the signal shame, the smaller is the effect.

Similarly, the relative Gaussian ( of the electron peak in the total reconstructed energy specrum for 4H/200 GeV/c runs (see Section 6 for details) is, respectively, 9.1% , 6.4% and 5.8%.

The spline method is used for all the numbers quoted further in this note.

3. Pedestals

We considered three methods to obtain pedestals for each cell:

- P1 – use A0  as a pedestal in each individual event;

- P2 –  use A0 averaged over beam events in the current run;

- P3 –  use pedestal events in the current run to obtain the average pedestal.

As seen from the Table 2, the method P2 gives a better energy resolution than the method P1. The methods P3 and  P2 give almost identical results.

“P2 “ pedestals are further used.

| Table 2: Measured energy resolution obtained with different pedestal subtraction methods |
| --- |
| Beam | P1 | P2 |
|  | (,% | (,% |
| E, 40GeV/c, 4H | 17.2 | 12.9 |
| E, 100GeV/c, 4H | 14.8 | 11.0 |
| E, 200GeV/c, 4H | 6.4 | 5.8 |
| , 200GeV/c, 4H | 14.4 | 11.4 |

4. Threatment of the “unsummed” cells

Signals from the unsummed cells are scaled down by factor two before adding them to integral quantities like the total response or the total noise. Thus, while a regular (“summed”) cell typically contributes an rms noise of about 3.2 ADC channels to a sum of cells, the contribution of an “unsummed” cell is about 1.6 ADC channels.

5. The calorometer response

The calorimeter response R is computed by summing all the cells together (for electrons ( in FEBs 6 and 7, for pions ( in FEBs 0, 1, 5, 6 and 7) or by selecting only the cells  within the radius r_core = 4, 8, 12, 16 cm from the cell with the highest amplitude. We refer to these R-values as the  total and core  responses, respectively. For pions, we apply relative weights to FCAL2 and FCAL3:

Re =  FCAL1 (FEB#6,7)

R( =  FCAL1 (FEB#6,7) + g2(FCAL2 (FEB#5)+g3(FCAL3 (FEB#0, 1)

The intercalibration factors g2 and g3 are obtained by minimizing the resulting energy resolution (see Fig.3). Data favor g2=g3=2. This value is in a good agreement with g2=2.1 resulting from 1998 calibrations of FCAL1 and FCAL2 with electrons.

Technically, the cell summation is performed as follows:

- first, the sample amplitudes Ai  in the selected cells are summed up, sample-by-sample, to obtain the sequence of 7 integrated sample amplitudes ;

- the filtering method (a spline fit in our case) is applied to this sequence.  In the present analysis, we abandoned the alternative approach, in which the  individual cell signals are first filtered and then the resulting amplitudes are summed, because of the ambiguities and fit instabilities in cells with very small or no signals.

- Event selection

Muon counter proved to be  the most effective tool to clean the data samples.  Cuts on other counters and BPC's  do not improve the situation (Fig.4a-4d, data - electrons, 4H, 200GeV/c), as Table 3 shows.

# 7. Electrons

The distributions of the calorimeter response for different core radii are shown in Fig.5. The resilts of the Gaussian fits of the electron peak are summarized in Table 4.

| Table 3: Gaussian mean total response and the resolution, as function of event selection cuts |
| --- |
| Cuts# | cut description | Number of events | , ADC counts | , % |
|  | no cuts | 10561 | 2324 | 5.8 |
|  | PH(muon counter)<20 | 7872 | 2324 | 5.8 |
|  | Cut#1 & PH(veto counter)<10 | 5864 | 2327 | 5.8 |
|  | Cut#1 & 0.5<counters_signal<1.5 | 5980 | 2329 | 5.9 |
|  | Cut#1 & 0.5<chambers_signal<1.54 | 5791 | 2328 | 5.8 |
|  | Cut#1 & ¦angleX¦<0.5mrad | 7117 | 2326 | 5.8 |
|  | Cut#1 & -1mrad<angleY<-0.5mrad | 7008 | 2329 | 5.8 |

| Table 4: Gaussian mean total response and the resolution, as function of r_core cuts, for electrons |
| --- |
| P, GeV/c | position | , ADC counts / , % |
|  |  | Total response | r_core, cm |
|  |  |  | 4 | 8 | 12 | 16 |
| 40 | 4H | 447 ¦ 12.9 | 423 ¦ 7.4 | 438 ¦ 7.4 | 443 ¦ 8.8 | 445 ¦ 10.5 |
| 100 | 4H | 1155 ¦ 11.0 | 1096 ¦ 5.5 | 1.131 ¦ 5.3 | 1143 ¦ 5.9 | 1148 ¦ 7.5 |
|  | 4L | 1161 ¦ 10.9 | 1107 ¦ 5.4 | 1139 ¦ 5.3 | 1151 ¦ 5.8 | 1153 ¦ 7.0 |
| 200 | 4H | 2329 ¦ 5.8 | 2237 ¦ 6.4 | 2302 ¦ 5.6 | 2324 ¦ 5.5 | 2327 ¦ 5.7 |
|  | 4L | 2196 ¦ 6.4 | 2116 ¦ 6.8 | 2168 ¦ 6.3 | 2188 ¦ 6.2 | 2193 ¦ 6.2 |

Observations:

- 99% of R_total is contained in a region with r_core<8cm. The Gaussian resolution for this region can be approximated by

- The  resolution for R_total  at 100 GeV/c is much too large. This results from a larger noise in runs at 100GeV/c, due to the pulser board noise. See Section 9 for more discussion.

	he R distributions are well described by Gaussian for one and half order of magnitude (Fig.6). At a lower level there is a long tail of smaller responses, which is partly due to beam pion contamination. The radial cuts reduce the tail (Fig.6-7).

  - A map of energy deposited  by electrons hitting the same cell in FCAL1  is shown in Fig.8.  This plot represents an electromagnetic shower profile.

# 8. Hadrons

Distributions of R and the Gaussian fit results for pion runs are shown in Fig.8 and in Table 6.

# Observations:

# The region of r_core<16cm contains ~95% of R_total.

# The energy resolution plotted as function of   is shown in Fig.9. The electronic noise is subtracted.  The stochastic term appears to be  , the constant term – (3-5)%.  The limited range of beam energies does not permit to make more precise estimates.

- Like for electrons, there is a pronounced noise problem at 100GeV/c, see Section 9.

| Table 5: Gaussian mean total response and the resolution, as function of r_core cuts, for pions |
| --- |
| P, GeV/c | position | , ADC counts / , % |
|  |  | Total response | r_core, cm |
|  |  |  | 4 | 8 | 12 | 16 |
| 60 | 4H | 567 ¦ 28.3 | 337 ¦ 37.5 | 459 ¦ 24.3 | 512 ¦ 20.7 | 540 ¦ 20.8 |
|  | 4L | 589 ¦ 26.4 | 362 ¦ 30.3 | 468 ¦ 20.0 | 522 ¦ 17.5 | 548 ¦ 17.4 |
| 100 | 4H | 991 ¦ 32.4 | 540 ¦ 35.7 | 755 ¦ 19.6 | 843 ¦ 17.5 | 889 ¦ 19.3 |
|  | 4L | 1033 ¦ 33.2 | 593 ¦ 25.0 | 772 ¦ 16.2 | 865 ¦ 15.3 | 915 ¦ 18.3 |
| 200 | 4H | 1921 ¦ 11.4 | 1267 ¦ 27.5 | 1654 ¦ 16.6 | 1803 ¦ 12.1 | 1872 ¦ 10.9 |
|  | 4L | 1894 ¦ 11.2 | 1305 ¦ 22.8 | 1627 ¦ 13.4 | 1755 ¦ 10.6 | 1841 ¦ 9.8 |

# 9. Electronic channel noise

The Table 6 shows the total rms noise in individual FEB's, evaluated by summing up all cells in pedestal events (with the “unsummed” cells scaled-down by factor 2!).  Table 7 shows the noise  contribution to the total response, different for electrons and pions because R(    and  Re  are composed of different combinations of the FEBs. Note, that the total noise does not exactly scale up from individual FEBs, as one would expect for totally independent noise sources (this is clearly seen for FEBs 6 and 7, which are summed-up in Re). Apparently, this is a manifestation of the coherent component of the electronic noise.

| Table 6: Total rms noise of the electronic channels in individual FEBs |
| --- |
| runs, 4H | FEB#0 | FEB#1 | FEB#5 | FEB#6 | FEB#7 |
| e, 40GeV/c | 50.9 | 47.8 | 50.8 | 41.5 | 43.1 |
| e, 100GeV/c | 57.0 | 66.5 | 63.4 | 65.4 | 60.7 |
| e, 200GeV/c | 48.3 | 49.2 | 50.8 | 41.2 | 43.4 |

| Table 7: the contribution of the electronics noise to Re and R( |
| --- |
| Runs: | e, 40, 4H | e, 100, 4H | e, 200, 4H |
| Total noise in contribution to Re FEBs 6 and 7 | 51.8 | 105.0 | 51.9 |
| Runs: | , 60, 4H | , 100, 4H | , 200, 4H |
| Total noise in contribution to R( FEBs 7,6,5,1 and 0 | 86.0 | 200.7 | 86.0 |

Observations:

- The individual channel noise for different FEB's  is plotted in Fig.10. There is an additional contribution in the first thirty channels for all runs at 100 GeV/c. One can see a correlation between the FEB's (Fig.11). In 40, 60, 200GeV/c runs that kind a noise is not observed.

- This effect is associated with the pulser board that was left enabled during the physics data taking for a large fraction of 100 GeV/c runs, in order to take pulser triggers along with beam triggers. In that mode, the pulser board was demonstrated to induce a strong cross-talk to the inputs of all FEBs. The origin of the noise inside the pulser board itself remained unclear, though.

- The step-like patterns in the pedestal distributions in Fig.10 correspond to groups of “unsummed” channels, in which the signal (hence, the noise) is scaled down by two.

# Figures

Fig.1. Example of a signal shape, electrons, 200GeV/c

Fig.2. Reconstructed amplitude distributions, for simulated data and different filtering methods.

Fig.3. The effect of the intercalibration coefficients g2, g3 on the response and the resolution for pions at 200GeV/c.  R=FCAL1+g2*FCAL2+g3*FCAL3

Fig.4a. Event selection, 4H, 200GeV/c,

___no cuts, ___PH(muon counter)<15, ___Gauss fit

Fig.4b. Event selection, 4H, 200GeV/c,

___PH(muon counter)<15, ___& PH(veto counter)<15, ___Gauss fit

Fig.4c-h. Plots R_total vs.:

c – counters_signal, d – chambers_signal,

e – angleX, f – angleY,

g – X of cryostat, h – Y of cryostat.

Electrons, 4H, 200GeV/c,

Fig.5. The distributions of  R_core(8cm), electrons, 4H:

a – 40GeV/c, b – 100GeV/c, c – 200GeV/c

Fig.6. The distribution of  R_core(4cm)/R_total, electrons, 4H, 200GeV/c

Fig.7. The distributions of  R_core(8cm):

___data, ___data(cut), ___Gauss,

cut: 0.9<R_core(4cm)/R_total<1.05

Fig.8. FCAL1 map, electrons, 4H, 200GeV/c:

+ - beam position

o – cell centers,

o – cells with 4 tubes,

. – cells with “negative” energy R,

___ - energy deposited inside a contour,

Min(R)=-0.15 ADC count. Total “negative” R=-1.8 ADC count (-0.08%R).

Fig.8. Distribution of R_core(16cm),

pions, 4H, a – 60GeV/c, b – 100GeV/c, c – 200GeV/c

Fig.9. Sigma(R_core(16cm)) vs.   for pions, 4H, with

the electronic channel noise  quadratically subtracted

Fig.10. Channel noise in different  FEBs,  in ADC counts

Рис.11. An example of a correlated noise in FEB6 and FEB7 (channel 20) , , 4H, 100GeV/c

# Footnotes

With the advent of  new high quality data from the September run, the results on electrons presented in this note have become somewhat obsolete. We have already circulated the summary plots and new estimates of the electron response and resolution  based on new 10-100 GeV/c data. However,  the analysis methods remained exactly the same. The changes compared to the original  draft version of this note are mostly stylistic and aimed at better description of the methods.

We found too few electrons in runs 3242, 3243, 3328, 3329 (60GeV/c). Therefore, the 60 GeV/c point for electrons is not presented in this note.

A triangular primary signal shape with t=50ns without fluctuation of an amplitude is used.

The quoted numbers correspond to the “Total response” columns of Tables 4 and 5.

;

# Text boxes

sigma(R), %

sigma(R), %

e, 100GeV/c, run2967

mean(R), ADC count

mean(R), ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_total, ADC count

R_core, ADC count

R_core, ADC count

R_core, ADC count

R_core(4cm)/R_total

R_core, ADC count

R_core(16cm), ADC count

R_core(16cm), ADC count

R_core(16cm), ADC count