V. Epshteyn (ITEP) , P.Gorbunov (UoT)

ATLAS-FCAL beam tests 2003 Internal Note 8, 2004-Feb-04; Updates: PG 10-12/02/2004

Some results of the amplifier performance modeling in the full range of input signal amplitudes

Abstract:  the FEB amplifier performance is simulated for the input signals of up to 5 TeV.  With the input cables connected, the pulser data is predicted to reproduce the amplifier response to real calorimeter signals up to 1 TeV. With a simple hardware modification of the pulser, one can reduce the non-linearity effects observed when the input cables are disconnected.

In connection with the input cables on/off problem [Ref.1], John Rutherfoord suggested to extend the model study to investigate

- a calibration with DAC (10000 for cables on | off;

- a possibility of using the measurements done with cables off to evaluate the input impedance of the amplifiers.

It is also interesting to see what the model predicts for a non-linearity in the entire range of the input physical signals.

The schematics for the model are given in [Ref.1], Figures 16 and 17.

Physics signal

Fig. 1 (blue line) shows the result for the case of physics signals in the range 10 GeV – 5 TeV.

An integral non-linearity of up to 2% is predicted by the model. Note that in real life there will be no signals of the order of 1 TeV in a single cell, because of the shower lateral spread. Hence, a typical response non-linearity will be smaller than 2%.

Calibration pulser signal, input cables on

Fig. 1 (red line) shows the model prediction for pulser calibration with the input cables connected.

In a wide range of input amplitudes (up to 1 TeV), the observed non-linearity is the same, within a fraction of a percent, for pulser and physics inputs. Therefore, using the pulser data as a correction, one can reduce the response non-linearity in case of physics signals.

Calibration pulser signal,  input cables off

As Fig. 2 shows, the non-linearity further keeps increasing for higher DAC. Model parameters Cpar and track length are explained in [Ref.1].

A possibility of using measurements with input cables off for estimation of the input impedance of the amplifiers. Two options can be considered:

- the integral non-linearity for DAC<1000 does not exceed 2%. This is predicted by the model and confirmed by data (Fig. 3). Therefore, one can use the measurements with DAC<1000 for the impedance estimates;

- as an alternative, one can change the pulser signal shape. The modification is very simple: a capacitor of ~300 pF has to be added in parallel to the 1.8 H inductance (Fig. 4). The problem with cables off largely disappears (Fig. 5).

References

1.  V. Epshteyn, P. Gorbunov, “A study of non-linear effects in FCAL pulser calibration data”, ATLAS-FCAL beam tests 2003 Internal Note 5, 2003-12-21, http://atlas-fcaltb.web.cern.ch/atlas-fcaltb/Memos/ (Analysis section).

Figures

Figure 1: Signal amplitude after the shaper, as function of the energy deposited in one FCAL-1 cel.l

The electromagnetic calorimeter - 1GeV ( 1.5A, 100GeV ( DAC=1000. Signals:

___physical - triangle, t=50ns,I=15A(8mA,

___calibration - input cables on, DAC=100(60000.

Figure 2: Signal amplitude after the shaper, as function of the pulser DAC, with input cables off and different parasitic capacitance (Cpar) and FEB PCB track length.

Figure 3:  The non-linearity in the calorimeter channels for DAC=250, with high gain, input cables off. DAC_max=1000.

Figure 4: The pulser signal shape, with and without modification (see text), for Cpar=2 pF.

The waveform is observed at the point where the pulser and the physical signals meet: before the T2 delay ([Ref.1], Fig. 16).  To exclude the influence of the pre-amplifier, T2 is set to 200 ns.

Figure 5:  Signal amplitude after shaper, with cables off and a  modified calibration pulse shape.   Сpar= 2pF, track: …23mm, ___88mm.

# Text boxes

Cpar=1 pf,  23 mm track

___original, ___modified,

track: …23 mm, ___88 mm

Cpar=2 pf,  23 mm track

Cpar=2 pf,  88 mm track

Cpar=1 pf,  88 mm track