3.1 Principle:

The calibration pulses are generated by pulsers the

principle of which is described on Figure 1.

The fast output voltage pulse is obtained by

interrupting a precise DC current Ip that flows in the

inductor. When a pulse command is applied on Q2, Q1 is

cut off and the current is diverted to ground. The

magnetic energy stored in the inductor produces a voltage

pulse with an exponential decay across the parallel

combination of the cable characteristic impedance Zc and

a termination resistor Ro of the same value. This pulse is

propagated inside the cryostat through a 7 m long 50 

cable adapted at both ends and is applied across a precise

injection resistor Rinj (0.1%) in the cold, directly to the

electrodes. The resulting current is given by:

An amplitude up to 10 mA can be achieved by

applying a 5V pulse in a 500 injection resistor.

The main difficulty of this calibration system is to

distribute uniformly throughout the calorimeter a very

precise voltage pulse. To minimise the attenuation due to

skin effect, the pulsers are located close to the

feedthrough.

Figure 1 : Principle of the pulser

3.2 Complete design

The current Ip is generated from a 18 bits DAC voltage

through a voltage to current converter. The current in the

output branch is I1=VDAC/R20, as negligible current

flows in the gate of the JFET transistor. A low offset op

amp as OP 07 must be used, as Voffset = 100 V already

corresponds to 15 LSB.

To insure good linearity, the base current of Q1 which

varies non linearly with Ip (varies with the collector

current and the temperature) is measured and added to Ip.

Q1 is made of 4 transistors in parallel to share the large

DC current for Ip (up to 200 mA) and to keep a low Rcc,

necessary for the wide dynamic range.

Five precision resistors (0.1%) determine the calibration pulse accuracy.

Figure 2 : Schematics of the analog part